## Slide 1: Innovation of Acoustic Research on Biomedical Applications
Jun Qin, Ph. D.
Assistant Professor
Department of Electrical and Computer Engineering The College of Engineering
Southern Illinois University, Carbondale, IL, USA

## Slide 2: Acoustic Research Projects in Our Lab
1. Applications on Therapeutic Ultrasound
--- Innovation of Shock Wave lithotripsy (SWL) on Treatment of Kidney Stone Diseases
--- Cavitation Bubbles – Cell Interaction for Ultrasound Enhanced Gene Activation
2. Applications on Audible Sound
--- Research on Noise-Induced Human Hearing Loss
--- Diagnosis and Treatment of Human Tinnitus

## Slide 3: PART I:
Comparison of Electrohydraulic (EH) and Electromagnetic (EM) SWLs
Introduction to SWLs.
Characterization of acoustic fields
Stone fragmentation in vitro and in vivo

## Slide 4
Shock Wave lithotripsy
P+
Lithotripter shock wave
P-

## Slide 5
Electromagnetic (EM) Shock Wave Generator
Widely use in the newer generation lithotripters
Stable and highly reproducible shock waves, long life time
High peak pressure and narrow focal beam size
1st generation---
Dornier HM-3:
“Gold standard”
Newer generation machines:
Less effective in stone comminution
Higher propensity for tissue injury
Newer is not better! Why?

## Slide 6
Comparison of EH and EM Lithotripters
Electrohydraulic:
Unmodified HM-3 at 20 kV
Electromagnetic
Siemens Modularis at E4.0
In vitro comparison
Acoustic fields Stone fragmentation
In vivo comparison
Stone fragmentation

## Slide 7: Acoustic Field Measurement
Light Spot Hydrophone (LSHD-2)
(Siemens/University of Erlangen-Nuremberg)

## Slide 8: Pressure Waveforms at Focus
HM3 at 20 kV
Dual-Peak Structure 1st P+
P-
Modularis at E4.0
2nd P+
1st
P+
P-

## Slide 9
Pressure Distribution and Characteristics of Acoustic Fields
Peak P+ Peak P- -6 dB Beam Size, -6dB Beam Size, Effective (MPa) (MPa) Head-Foot (mm) Left-Right (mm) Energy (mJ) |  |  |  |  | 
HM3 at 20 kV | 48.9  1.3 | -10.7  0.6 | 12.5 | 9.3 | 42.9
Modularis at E4.0 | 54.3  1.0 | -14.4  3.4 | 6.8 | 6.6 | 62.1

## Slide 10: Stone Fragmentation in a Finger Cot Holder
15mm
Stone fragments are always kept in a 15 mm diameter area during SWL
Do not represent stone fragmentation in vivo
Finger Cot Holder

## Slide 11: New Stone Holder: Membrane Holder
2000 shocks in vivo
30 mm
15mm
Focal Area of HM-3
Stone fragments are always kept in a 15 mm diameter area during SWL
Do not represent stone fragmentation in vivo
Finger Cot Holder
Allow fragments to accumulate & spread out laterally
Mimic more closely stone fragmentation in vivo
30 mm
Membrane Holder

## Slide 12
Spreading of Fragments in a Membrane Holder
Focal Area of HM-3
0 shocks
50 shocks
100 shocks
250 shocks
500 shocks
1000 shocks
1500 shocks
2000 shocks

## Slide 13: Stone Fragmentation in a Member Holder
Allow fragments to accumulate & spread out laterally
Mimic more closely stone fragmentation in vivo
30 mm
Membrane Holder

## Slide 14
Stone Fragmentation in vivo

## Slide 15: Summary of PART I Stone fragmentation produced by the EM lithotripter is lower than that of the EH lithotripter both in vivo and in the membrane holder.
The acoustic field characterization demonstrates two distinct differences between EM and EH lithotripters
  2nd compressive component in EM pulse, which could reduce maximum bubble size by 50%
  EM lithotripter has much narrower beam size than EH lithotripter

## Slide 16: PART II:
Development of a Noise Exposure System for Research on Impulse Noise Induced Hearing Loss
Anatomy of Human Ear
Noise Induced Hearing Loss
Development of the Impulse Noise Exposure System

## Slide 17: Noise-Induced Hearing Loss (NIHL)
When an individuals hearing is damaged or altered by noise
  One of the most common occupational disabilities in the United States
    More than 30,000,000 American workers exposed to unsafe noise levels at their job
    Estimated 600 million people worldwide exposed to hazardous noise levels

## Slide 18: Anatomy of the Human Ear
Pinna
Tragus
Exterior Auditory Canal
Tympanic Membrane
Ossicles
Scala Vestibuli
Scala Tympani
Cochlea

## Slide 19
Anatomical Areas Affected by Different Noise Exposure
Gaussian Noise (Steady State Noise): Cochlea and Stria Vascularis
Impulse/Impact Noise: Cochlea and Stria Vascularis and possible tympanic membrane and ossicular damage depending on level
Kurtosis Noise (complex noise): A combination of Gaussian and impulse/impact noise which can damage all of the above areas depending on the noise content.

## Slide 20: Noise Exposure System

## Slide 21: Impulse Noise
(http://www.arl.army.mil/www/default.cfm?page=352)
Friedlander’s Wave
(1 t * )
P(t)  Ps e
t / t*
t
Ps = peak sound pressure
t* = the time at which the pressure crosses the x-axis

## Slide 22
3
Time (ms)
0
1
2
4
5
6
Pressure (Pa)
0
-500
500
1000
1500
Peak SPL = 158 dB
Simulated Wave vs. Field-Measured Wave
4
6
8
10
×10
-3
TIME IN SECONDS
P
E
R
S
S
U
R
E
0
-200
-100
100
0
200
400
300

## Slide 23
Peak Pressure vs. Output Voltage
dB
Pref
P(t)
SPL  20 log
10

## Slide 24
Animal Study Verifying the Impulse
Noise Induced Hearing Loss
Animal Model: Chinchilla
  10150daBniSmPaLls were tested
Noise Exposure: impulse noise with peak SPL=155 dB at 2 Hz pulse repartition rate for 75 seconds (150 pulses).
Auditory brainstem response (ABR) were measured before and 21 days after noise exposure

## Slide 25: Animal Study Results
105 dB SPL

## Slide 26: Summary of PART II
A digital noise exposure system has been developed to generate the impulse noise.
The waveforms of impulse noise are comparable to the field measurement test performed by the U.S. Army
Impulse noise produces significant hearing loss in animal study.
Future work includes Kurtosis noise simulate and high level impulse noise generation.

## Slide 27: Upcoming Conference
For upcoming conferences please follow the below mentioned link
http://www.conferenceseries.com/